## 5.1
Test: | 5.1
Comparison: | Dose distributions in planning module vs. modeling (physics) module
Description: | Comparison of dose distribution for large (>30x30) field.
Tolerance: | Identical
Comments: | Our validation method includes a direct comparison between calculated data in planning mode and measured data from commissioning. Therefore this test in unnecessary.
Results: | See Test 5.3

## 5.2
Test: | 5.2
Comparison: | D/MU in test plan vs. clinical calibration condition
Description: | Reference calibration condition check
Tolerance: | 0.5%
Comments: | Our calibration technique involves setting the dose to 1 cGy/MU at dmax for 100 SSD setup. There are two ways to assess this. One is to determine the dose per MU at dmax and verify that it is 1 cGy/MU. Another would be to determine the dose/MU at 10 cm, since this is where the calibration measurement is performed. Below, we verify the dose/MU at dmax and the PDD at 10 cm.
Test Patient: | zz MPPG #5 Tests
Test Course: | MPPG Photons
Test Plan: | 5.2 TB140
Plan Settings: | 2 mm dose grid
Results:
Eclipse AAA
|  | TPS Data |  |  | Dose/MU at Dmax |  |  | Percent Depth Dose at Reference Depth
Energy | Ref. Depth | MU | Ref. Dose [cGy] | MU/Gy (10 cm) | Dose/MU (REF) | % Diff. | Within 0.5%? | PDD10 (Eclipse) | PDD10 (Databook) | % Diff. | Within 0.5%?
6x | 1.4 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 66.3% | #DIV/0! | #DIV/0!
6x FFF | 1.3 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 63% | #DIV/0! | #DIV/0!
10x FFF | 2.2 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 70.5% | #DIV/0! | #DIV/0!
16x | 2.9 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 76.5% | #DIV/0! | #DIV/0!
Eclipse Acuros
|  | TPS Data |  |  | Dose/MU at Dmax |  |  | Percent Depth Dose at Reference Depth
Energy | Ref. Depth | MU | Ref. Dose [cGy] | MU/Gy (10 cm) | Dose/MU (REF) | % Diff. | Within 0.5%? | PDD10 (Eclipse) | PDD10 (Databook) | % Diff. | Within 0.5%?
6x | 1.4 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 66.3% | #DIV/0! | #DIV/0!
6x FFF | 1.3 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 63% | #DIV/0! | #DIV/0!
10x FFF | 2.2 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 70.5% | #DIV/0! | #DIV/0!
16x | 2.9 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 76.5% | #DIV/0! | #DIV/0!

## 5.3
Test: | 5.3
Comparison: | Dose distribution calculated in planning system vs. commissioning data
Description: | Comparison of dose distribution for large (>30x30) field.
Tolerance: | Identical
Comments: | This analysis is performed using Nick's MATLAB code. Results are stored in PDFs. Links to these PDFs are found below.
Test Patient:
Test Course:
Test Plan:
Plan Settings:
Results: | See PDFs

## 5.4
Test: | 5.4
Description: | Small MLC shaped field (non SRS)
Comments: | The field shape for this test is shown to the right.
Test Patient: | zz MPPG #5 Tests
Test Course: | MPPG Photons
Test Plan: | 5.4 TB140
Plan Settings: | 2 mm dose grid, all beams have 500 MU
Scan SSD: | 90 cm
Crossline Profile Depths: | dmax, 10 cm, 25 cm
Inline Profile Depths: | 10 cm
Point Dose Results:
|  |  |  | Measurement |  |  |  |  | AAA |  |  | Acuros
Plan Name | Field Name | Dose Rate | Description | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff
5.2 TB140 | 6x_10x10 | 600 Servo On | 6x 10x10 Standard |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.2 TB140 | 6xFFF_10x10 | 800 Servo On | 6x FFF 10x10 Standard |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.2 TB140 | 10xFFF_10x10 | 800 Servo On | 10x FFF 10x10 Standard |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.2 TB140 | 16x_10x10 | 600 Servo On | 16x 10x10 Standard |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.4 8 mm Offset | 6x | 600 Servo On | MPPG 5.4 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.4 8 mm Offset | 6x FFF | 800 Servo On | MPPG 5.4 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.4 8 mm Offset | 10x FFF | 800 Servo On | MPPG 5.4 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.4 8 mm Offset | 16x | 600 Servo On | MPPG 5.4 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
|  | AAA |  |  |  |  | Acuros
Profile Passing Rates: |  | Crossline |  |  | Inline | PDD | Crossline |  |  | Inline | PDD
Criteria: 2%/2mm Global |  | dmax | 10 cm | 30 cm | 10 cm |  | dmax | 10 cm | 30 cm | 10 cm
5.4 8 mm Offset | 6x
5.4 8 mm Offset | 6x FFF
5.4 8 mm Offset | 10x FFF
5.4 8 mm Offset | 16x

## 5.5
Test: | 5.5
Description: | Large MLC shaped field with extensive blocking (e.g.: mantle)
Comments: | The field shape for this test is shown to the right.
Test Patient: | zz MPPG #5 Tests
Test Course: | MPPG Photons
Test Plan: | 5.5 TB140
Plan Settings: | 2 mm dose grid, all beams have 500 MU
Scan SSD: | 90 cm
Crossline Profile Depths: | dmax, 10 cm, 25 cm
Inline Profile Depths: | 10 cm
Point Dose Results:
|  |  |  | Measurement |  |  |  |  | AAA |  |  | Acuros
Plan Name | Field Name | Dose Rate | Description | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff
5.2 TB140 | 6x_10x10 | 600 Servo On | 6x 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 6xFFF_10x10 | 800 Servo On | 6x FFF 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 10xFFF_10x10 | 800 Servo On | 10x FFF 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 16x_10x10 | 600 Servo On | 16x 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.5 TB140 | 6x | 600 Servo On | MPPG 5.5 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.5 TB140 | 6x FFF | 800 Servo On | MPPG 5.5 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.5 TB140 | 10x FFF | 800 Servo On | MPPG 5.5 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.5 TB140 | 16x | 600 Servo On | MPPG 5.5 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
|  | AAA |  |  |  |  | Acuros
Profile Passing Rates: |  | Crossline |  |  | Inline | PDD | Crossline |  |  | Inline | PDD
Criteria: 2%/2mm Global |  | dmax | 10 cm | 30 cm | 10 cm |  | dmax | 10 cm | 30 cm | 10 cm
5.5 | 6x
5.5 | 6x FFF
5.5 | 10x FFF
5.5 | 16x

## 5.6
Test: | 5.6
Description: | Off-axis MLC shaped field, with maximum allowed leaf over travel.
Comments: | The field shape for this test is shown to the right.
Test Patient: | zz MPPG #5 Tests
Test Course: | MPPG Photons
Test Plan: | 5.6 TB140
Plan Settings: | 2 mm dose grid, all beams have 500 MU
Scan SSD: | 90 cm
Crossline Profile Location: | z = 12 cm
Crossline Profile Depths: | 10 cm
Inline Profile Location: | x = 10 cm
Inline Profile Depths: | dmax, 10 cm, 25 cm
Point Dose Results:
|  |  |  | Measurement |  |  |  |  | AAA |  |  | Acuros
Plan Name | Field Name | Dose Rate | Description | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff
5.2 TB140 | 6x_10x10 | 600 Servo On | 6x 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 6xFFF_10x10 | 800 Servo On | 6x FFF 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 10xFFF_10x10 | 800 Servo On | 10x FFF 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 16x_10x10 | 600 Servo On | 16x 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.6 TB140 | 6x | 600 Servo On | MPPG 5.6 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.6 TB140 | 6x FFF | 800 Servo On | MPPG 5.6 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.6 TB140 | 10x FFF | 800 Servo On | MPPG 5.6 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.6 TB140 | 16x | 600 Servo On | MPPG 5.6 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
|  | AAA |  |  |  |  | Acuros
Profile Passing Rates: |  | Inline |  |  | Crossline | PDD | Inline |  |  | Crossline | PDD
Criteria: 2%/2mm Global |  | dmax | 10 cm | 30 cm | 10 cm |  | dmax | 10 cm | 30 cm | 10 cm
5.6 | 6x
5.6 | 6x FFF
5.6 | 10x FFF
5.6 | 16x

## 5.7
Test: | 5.7
Description: | Asymmetric MLC shaped field at minimal anticipated SSD
Comments: | The field shape for this test is shown to the right.
Test Patient: | zz MPPG #5 Tests
Test Course: | MPPG Photons
Test Plan: | 5.7 TB140
Plan Settings: | 2 mm dose grid, all beams have 500 MU
Scan SSD: | 80 cm
Crossline Profile Location: | z = 0
Crossline Profile Depths: | dmax, 10 cm, 25 cm
Inline Profile Location: | x = 0
Inline Profile Depths: | 10 cm
Point Dose Results:
|  |  |  | Measurement |  |  |  |  | AAA |  |  | Acuros
Plan Name | Field Name | Dose Rate | Description | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff
5.2 TB140 | 6x_10x10 | 600 Servo On | 6x 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 6xFFF_10x10 | 800 Servo On | 6x FFF 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 10xFFF_10x10 | 800 Servo On | 10x FFF 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 16x_10x10 | 600 Servo On | 16x 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.7 TB140 | 6x | 600 Servo On | MPPG 5.7 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.7 TB140 | 6x FFF | 800 Servo On | MPPG 5.7 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.7 TB140 | 10x FFF | 800 Servo On | MPPG 5.7 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.7 TB140 | 16x | 600 Servo On | MPPG 5.7 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
|  | AAA |  |  |  |  | Acuros
Profile Passing Rates: |  | Crossline |  |  | Inline | PDD | Crossline |  |  | Inline | PDD
Criteria: 2%/2mm Global |  | dmax | 10 cm | 30 cm | 10 cm |  | dmax | 10 cm | 30 cm | 10 cm
5.7 | 6x
5.7 | 6x FFF
5.7 | 10x FFF
5.7 | 16x

## 5.8
Test: | 5.8
Description: | MLC shaped field at oblique incidence (30°)
Comments: | The field shape for this test is shown to the right.
Test Patient: | zz MPPG #5 Tests
Test Course: | MPPG Photons
Test Plan: | 5.8 TB140
Plan Settings: | 2 mm dose grid, all beams have 500 MU
Scan SSD: | 90 cm
Crossline Profile Location: | z = 0
Crossline Profile Depths: | dmax, 10 cm, 25 cm
Inline Profile Location: | x = 0
Inline Profile Depths: | 10 cm
Point Dose Results:
|  |  |  | Measurement |  |  |  |  | AAA |  |  | Acuros
Plan Name | Field Name | Dose Rate | Description | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff
5.2 TB140 | 6x_10x10 | 600 Servo On | 6x 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 6xFFF_10x10 | 800 Servo On | 6x FFF 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 10xFFF_10x10 | 800 Servo On | 10x FFF 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.2 TB140 | 16x_10x10 | 600 Servo On | 16x 10x10 Standard | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
5.8 TB140 | 6x | 600 Servo On | MPPG 5.8 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.8 TB140 | 6x FFF | 800 Servo On | MPPG 5.8 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.8 TB140 | 10x FFF | 800 Servo On | MPPG 5.8 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
5.8 TB140 | 16x | 600 Servo On | MPPG 5.8 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
|  | AAA | [5 empty] | Acuros
Profile Passing Rates: |  | Crossline |  |  | Inline | PDD | Diagonal | Crossline |  |  | Inline | PDD | Diagonal
Criteria: 2%/2mm Global |  | dmax | 10 cm | 30 cm | 10 cm |  |  | dmax | 10 cm | 30 cm | 10 cm
5.8 | 6x
5.8 | 6x FFF
5.8 | 10x FFF
5.8 | 16x

## 5.9
Test: | 5.9
Description: | Large (>15cm) MLC field for each a non-physical wedge angle
Comments: | The field shape for this test is shown to the right. Measured using MapCHECK2 at three depths in Solid Water using an isocentric geometry.
Test Patient: | zz MPPG #5 Tests
Test Course: | MPPG Photons
Test Plan: | 5.9 TB140 (05, 10, 20)
Plan Settings: | 2 mm dose grid, all beams have 300 MU
Energy | Angle | Direction | Depth | AAA | Acuros | (3%/2mm)
6x | 60 | In | 5 cm
6x | 60 | Out | 5 cm
16x | 60 | In | 5 cm
16x | 60 | Out | 5 cm
6x | 60 | In | 10 cm
6x | 60 | Out | 10 cm
6x | 45 | In | 10 cm
6x | 30 | In | 10 cm
6x | 25 | In | 10 cm
6x | 20 | In | 10 cm
6x | 15 | In | 10 cm
6x | 10 | In | 10 cm
16x | 60 | In | 10 cm
16x | 60 | Out | 10 cm
16x | 45 | In | 10 cm
16x | 30 | In | 10 cm
16x | 25 | In | 10 cm
16x | 20 | In | 10 cm
16x | 15 | In | 10 cm
16x | 10 | In | 10 cm
6x | 60 | In | 20 cm
6x | 60 | Out | 20 cm
16x | 60 | In | 20 cm
16x | 60 | Out | 20 cm

## 6.1
Test: | 6.1
Objective: | Validate planning system reported electron (or mass) densities against known values.
Description: | CT-density calibration for air, lung, water, dense bone, and possibly additional tissue types.
805169 Tomotherapy TE Rod Set Specifications
SN | 805169-1137
CT Scanner | Downtown 1st Floor
CT ID in Eclipse | Host 7114
MRN | QA-CHZ
Course | Electron Density
Plan | Downtown CT
Commissioning
Batch Number | Rod Type Substitute | Mean HU | Calibrated Electron Density Relative to Water | Eclipse Relative Electron Density | % Diff | Calibrated Physical Density (g/cc) | Eclipse Mass Density (g/cc) | % Diff | AXB Reported Material
NA | Air |  | 0.001 |  | -100% | 0.0012 |  | -100% | Air
15051-10 | LN-300 Lung |  | 0.312 |  | -100% | 0.32 |  | -100% | Lung
15002-14 | LN-450 Lung |  | 0.454 |  | -100% | 0.47 |  | -100% | Lung
14854 | AP6 Adipose |  | 0.925 |  | -100% | 0.942 |  | -100% | Adipose
14895 | BR-12 Breast |  | 0.956 |  | -100% | 0.979 |  | -100% | 69% Adipose, %31 Muscle Skeletal
4435 | CT Solid Water |  | 0.989 |  | -100% | 1.018 |  | -100% | Muscle Skeletal
14913 | BRN-SR2 Brain |  | 1.049 |  | -100% | 1.053 |  | -100% | Muscle Skeletal
15029 | LV1 Liver |  | 1.063 |  | -100% | 1.095 |  | -100% | Cartilage
14789 | IB Inner Bone |  | 1.092 |  | -100% | 1.139 |  | -100% | 91% Cartilage, 9% Bone
14988 | B200 Bone Mineral |  | 1.097 |  | -100% | 1.144 |  | -100% | 90% Cartilage, 10% Bone
14905 | CB2-30% CaCO3 |  | 1.279 |  | -100% | 1.334 |  | -100% | 53% Cartilage, 47% Bone
15899 | CB2-50% CaCO3 |  | 1.472 |  | -100% | 1.562 |  | -100% | 4% Cartilage, 96% Bone
14867 | SB3 Cortical Bone |  | 1.696 |  | -100% | 1.824 |  | -100% | Bone

## 6.2
Test: | 6.2
Objective: | Heterogeneity correction distal and proximal to lung tissue
Description: | 5x5 cm2, measure dose ratio above and below heterogeneity outside of the buildup region
Tolerances: | 3%
Comments: | This phantom consists of 5 cm of Solid Water on top, 13 cm of lung substitute material in the middle, then 10 cm of Solid Water on the bottom. Cork was used for a lung substitute material. Ion chamber measurements were taken at 4 cm depth and 22 cm depth. 90 SSD.
Point Dose Results:
|  |  |  | Measurement | [6 empty] | AAA |  |  | Acuros (No Override, Dm) |  |  | Acuros (No Override, Dw) |  |  | Acuros (Override)
Energy | Field Size | SSD | Depth | T | P | M1 | M2 | Mave | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff
6x | 10x10 | 100 | 1.4 |  |  |  | -37.2 | -37.2 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
6xFFF | 10x10 | 100 | 1.3 |  |  |  | -36.518 | -36.518 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
10xFFF | 10x10 | 100 | 2.2 |  |  |  | -36.739 | -36.739 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
16x | 10x10 | 100 | 2.9 |  |  |  | -37.53 | -37.53 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
6x | 5x5 | 90 | 4 |  |  |  | -38.612 | -38.612 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
6xFFF | 5x5 | 90 | 4 |  |  |  | -37.821 | -37.821 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
10xFFF | 5x5 | 90 | 4 |  |  |  | -41.008 | -41.008 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
16x | 5x5 | 90 | 4 |  |  |  | -42.097 | -42.097 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
6x | 5x5 | 90 | 22 |  |  |  | -18.814 | -18.814 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
6xFFF | 5x5 | 90 | 22 |  |  |  | -17.627 | -17.627 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
10xFFF | 5x5 | 90 | 22 |  |  |  | -21.305 | -21.305 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
16x | 5x5 | 90 | 22 |  |  |  | -23.467 | -23.467 | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!

## 7.1 & 7.2
Test: | 7.1 & 7.2
Objective: | Very small field PDD & Verify output for small MLC-defined field
Description: | 1x1 cm2 MLC shaped field. Jaws are retract 0.8 cm in X and 0.2 cm in Y, as is done automatically. Measured with Edge detector.
Test Patient: | zz MPPG #5 Tests
Test Course: | MPPG Photons
Test Plan: | 7.1 TB140
Plan Settings: | 1 mm dose grid, all beams have 500 MU
Scan SSD: | 90 cm
Crossline Profile Depths: | dmax, 10 cm, 25 cm
Inline Profile Depths: | 10 cm
Point Dose Results: | [11 empty] | With kΩ = 0.967 applied to 1x1 fields
|  |  |  | Measurements |  | AAA |  |  | Acuros |  |  | Measurements |  | AAA |  |  | Acuros
Plan Name | Field Name | Dose Rate | Description | M | Rel. Dose | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff | M | Rel. Dose | Calc Dose. | Rel. Dose | % Diff | Calc Dose. | Rel. Dose | % Diff
7.0 TB140 | 7.0-6x | 600 Servo On | 6x 10x10 90 SSD |  | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | 0 | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
7.0 TB140 | 7.0-6xFFF | 800 Servo On | 6x FFF 10x10 90 SSD |  | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | 0 | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
7.0 TB140 | 7.0-10xFFF | 800 Servo On | 10x FFF 10x10 90 SSD |  | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | 0 | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
7.0 TB140 | 7.0-16x | 600 Servo On | 16x 10x10 90 SSD |  | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | 0 | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
7.1 TB140 | 7.1-6x | 600 Servo On | MPPG 7.1 |  | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | 0 | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
7.1 TB140 | 7.1-6xFFF | 800 Servo On | MPPG 7.1 |  | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | 0 | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
7.1 TB140 | 7.1-10xFFF | 800 Servo On | MPPG 7.1 |  | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | 0 | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
7.1 TB140 | 7.1-16x | 600 Servo On | MPPG 7.1 |  | #DIV/0! |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | 0 | #DIV/0! | 0 | #DIV/0! | #DIV/0! | 0 | #DIV/0! | #DIV/0!
|  | AAA |  |  |  |  | Acuros
Profile Passing Rates: |  | Crossline |  |  | Inline | PDD | Crossline |  |  | Inline | PDD
Criteria: 2%/2mm Global |  | dmax | 10 cm | 30 cm | 10 cm |  | dmax | 10 cm | 30 cm | 10 cm
7.1 | 6x
7.1 | 6x FFF
7.1 | 10x FFF
7.1 | 16x
Comments: According to "Francescon, Cora, and Satariano: Correction factors for several detectors using MC simulation, Med Phys 2011", the Edge Detector exhibits a 3.0-3.5% overreponse at 6x for a field size of 1x1. The response for other energies is unknown, but it is likely the overresponse is greater for higher energies. If we apply a uniform 3.3% correction, all of our measured output factors for 1x1 would be within 3% of the calculated doses for both algorithms. The agreement may be even better, but it isn't possible to apply energy-dependant corrections because they have not been published and require Monte Carlo to calculate.

## 7.3
Test: | 7.3
Objective: | TG-119 Tests
Description: | Plan, measure and compare planning and QA results to the TG-119 plans
See the link below for TG119 results:
W:\Private\Physics\ACCELERATORS\Varian_TB140_MP\TG119\IMRT_VMAT Results Table .xlsx

## 7.4
Test: | 7.4
Objective: | Clinical Tests
Description: | Choose at least 2 relevant clinical cases. Plan, measure, and perform an in-depth analysis of the results.
See:
Cheese Phantom\Film and IC\MPPG 7.4 Measurements.xlsx
W:\Private\Physics\ACCELERATORS\Varian_TB140_MP\MPPG 5 Testing\Cheese Phantom\Results\Summary PPT.pptx

## 8.0
Test: | 8.0
Comparison: | D/MU in test plan vs. clinical calibration condition for electrons
Description: | Reference calibration condition check
Tolerance: | 0.5%
Comments: | Our calibration technique involves setting the dose to 1 cGy/MU at dmax for 100 SSD setup and the 15x15 open cone.
Test Patient:
Test Course: | MPPG
Test Plan: | Standards
Plan Settings: | 2 mm dose grid
Results:
Eclipse Electron Monte Carlo
Energy | MU | Dose at Nominal Dmax |  |  |  |  | Dose at TPS Dmax
|  | Nominal Dmax | Dose Dmax,nom | Dose/MU | % Diff. | Within 0.5%? | TPS Dmax | Dose Dmax,TPS | Dose/MU | % Diff. | Within 0.5%?
6e |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  |  | #DIV/0! | #DIV/0! | #DIV/0!
9e |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  |  | #DIV/0! | #DIV/0! | #DIV/0!
12e |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  |  | #DIV/0! | #DIV/0! | #DIV/0!
16e |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  |  | #DIV/0! | #DIV/0! | #DIV/0!
16e |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! |  |  | #DIV/0! | #DIV/0! | #DIV/0!

## 8.1
Test: | 8.1
Objective: | Basic model verification with shaped fields
Description: | Custom cutouts at standard and extended SSDs.
Tolerance: | 3%/3mm
Test Patient:
Test Course: | MPPG
Test Plan: | 8.1 100 SSD, 8.1 105 SSD
Plan Settings: | 2 mm dose grid, all beams have 500 MU
Scan SSD: | 100 cm, 105 cm
Crossline Profile Depths: | Energy Dependant
Inline Profile Depths: | Energy Dependant
Point Dose Results:
|  |  |  | Measurement |  |  |  |  | EMC
Plan Name | Field Name | Dose Rate | Description | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff
Standards | 06e | 600 Servo On | 6e 10x10 Standard |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Standards | 09e | 600 Servo On | 9e 10x10 Standard |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Standards | 12e | 600 Servo On | 12e 10x10 Standard |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Standards | 16e | 600 Servo On | 16e 10x10 Standard |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Standards | 20e | 600 Servo On | 20e 10x10 Standard |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 100 SSD | 06e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 100 SSD | 09e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 100 SSD | 12e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 100 SSD | 16e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 100 SSD | 20e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 105 SSD | 06e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 105 SSD | 09e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 105 SSD | 12e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 105 SSD | 16e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.1 105 SSD | 20e | 600 Servo On | MPPG 8.1 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Electron Monte Carlo
Profile Passing Rates: |  | Crossline |  | Inline |  | PDD
Criteria: 3%/3mm Global |  | Depth 1 | Depth 2 | Depth 1 | Depth 2
8.1 100 SSD | 06e
8.1 100 SSD | 09e
8.1 100 SSD | 12e
8.1 100 SSD | 16e
8.1 100 SSD | 20e
Electron Monte Carlo
Profile Passing Rates: |  | Crossline |  | Inline |  | PDD
Criteria: 3%/3mm Global |  | Depth 1 | Depth 2 | Depth 1 | Depth 2
8.1 105 SSD | 06e
8.1 105 SSD | 09e
8.1 105 SSD | 12e
8.1 105 SSD | 16e
8.1 105 SSD | 20e

## 8.2
Test: | 8.2
Objective: | Surface irregularities - obliquity
Description: | Oblique incidence using reference cone and nominal clinical SSD. Measurement was done with 10x10 applicator, 105 SSD and 30 degree gantry rotation.
Tolerance: | 5%
Test Patient:
Test Course: | MPPG
Test Plan: | 8.2 105 SSD
Plan Settings: | 2 mm dose grid, all beams have 500 MU
Scan SSD: | 105 SSD
Crossline Profile Depths: | Energy Dependant
Inline Profile Depths: | Energy Dependant
Point Dose Results:
[7 empty] | Measurement |  |  |  |  | EMC
Plan Name | Field Name | Dose Rate | Description | X | Y | Z | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff
Standards | 06e | 600 Servo On | 6e 10x10 Standard | 0 | 0 | 1.25 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Standards | 09e | 600 Servo On | 9e 10x10 Standard | 0 | 0 | 2.05 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Standards | 12e | 600 Servo On | 12e 10x10 Standard | 0 | 0 | 2.8 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Standards | 16e | 600 Servo On | 16e 10x10 Standard | 0 | 0 | 3.3 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Standards | 20e | 600 Servo On | 20e 10x10 Standard | 0 | 0 | 2.6 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.2 105 SSD | 06e | 600 Servo On | MPPG 8.1 | 3.13 | 0 | 1.08 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.2 105 SSD | 09e | 600 Servo On | MPPG 8.1 | 3.53 | 0 | 1.78 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.2 105 SSD | 12e | 600 Servo On | MPPG 8.1 | 3.9 | 0 | 2.42 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.2 105 SSD | 16e | 600 Servo On | MPPG 8.1 | 4.15 | 0 | 2.86 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
8.2 105 SSD | 20e | 600 Servo On | MPPG 8.1 | 3.8 | 0 | 2.25 |  |  |  | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0!
Electron Monte Carlo
Profile Passing Rates: |  | Crossline |  | Inline |  | PDD | Diagonal
Criteria: 3%/3mm Global |  | Depth 1 | Depth 2 | Depth 1 | Depth 2
8.2 105 SSD | 06e
8.2 105 SSD | 09e
8.2 105 SSD | 12e
8.2 105 SSD | 16e
8.2 105 SSD | 20e

## 8.3
Test: | 8.3
Objective: | Inhomogeneity test
Description: | Reference cone and nominal clinical SSD.
Tolerance: | 7%
Test Patient: | zz MPPG #5 Tests
Test Course: | Hetero Check
Test Plan: | 6 MeV Hetero, 9 MeV Hetero, 12 MeV Hetero
Plan Settings: | 2 mm dose grid, all beams have 200 MU
| Measurements in Reference Geometry | [5 empty] | Electron Monte Carlo
Energy | Measured Dmax | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff
6 MeV | 1.2 | 23.5 | 769 | -3.836 | -3.8104 | 1 | 198.5 | 1 | 0%
9 MeV | 2 | 23.4 | 769 | -3.862 | -3.8349 | 1 | 199.6 | 1 | 0%
12 MeV | 2.8 | 23.4 | 769 | -3.899 | -3.8716 | 1 | 199.5 | 1 | 0%
16 MeV | 3.3 | 23.4 | 769 | -3.968 | -3.9402 | 1 | 199.9 | 1 | 0%
20 MeV | 2.6 | 23.4 | 769 | -4.075 | -4.0464 | 1 | 199.1 | 1 | 0%
| Measurements near Dmax in Heterogenity Phantom | [5 empty] | Electron Monte Carlo
Energy | Depth | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff
6 MeV | 2 | 23.4 | 768 | -3.749 | -3.7275 | 0.97826 | 191.6 | 0.96524 | -1.331%
9 MeV | 3.5 | 23.4 | 769 | -3.766 | -3.7396 | 0.97514 | 189.4 | 0.9489 | -2.691%
12 MeV | 5.5 | 23.4 | 769 | -3.657 | -3.6313 | 0.93793 | 182.8 | 0.91629 | -2.307%
16 MeV | 5.5 | 23.4 | 769 | -3.765 | -3.7386 | 0.94884 | 184.3 | 0.92196 | -2.833%
20 MeV | 5.5 | 23.4 | 769 | -3.847 | -3.82 | 0.94405 | 184.1 | 0.92466 | -2.054%
| Measurements In Longitudinal Dose Fall-off Region | [5 empty] | Electron Monte Carlo
Energy | Depth | T | P | M | Mcorr | Rel. Dose | Calc Dose. | Rel. Dose | % Diff
6 MeV | 3 | 23.4 | 768 | -2.053 | -2.0412 | 0.53571 | 97.2 | 0.48967 | -8.594%
6 MeV | 3.1 | 23.4 | 768 | -1.765 | -1.7549 | 0.46056 | 97.2 | 0.48967 | 6.321%
9 MeV | 5 | 23.4 | 769 | -1.893 | -1.8797 | 0.49016 | 100.6 | 0.50401 | 2.825%
9 MeV | 4.9 | 23.4 | 769 | -2.107 | -2.0922 | 0.54557 | 100.6 | 0.50401 | -7.618%
12 MeV | 8 | 23.4 | 769 | -1.663 | -1.6513 | 0.42652 | 83.9 | 0.42055 | -1.399%
12 MeV | 8.1 | 23.4 | 769 | -1.514 | -1.5034 | 0.3883 | 83.9 | 0.42055 | 8.304%
16 MeV | 9.5 | 23.4 | 769 | -1.858 | -1.845 | 0.46825 | 88.5 | 0.44272 | -5.451%
16 MeV | 9.6 | 23.4 | 769 | -1.747 | -1.7347 | 0.44027 | 88.5 | 0.44272 | 0.5563%
20 MeV | 11 | 23.4 | 769 | -1.975 | -1.9611 | 0.48466 | 94.2 | 0.47313 | -2.38%
20 MeV | 11.1 | 23.4 | 769 | -1.883 | -1.8698 | 0.46209 | 94.2 | 0.47313 | 2.39%
| Electron Monte Carlo |  |  | Depth at which Measured Ratio Equals Calculated Ratio |  |  | Distance to Agreement (mm)
Energy | Depth | Calc Dose. | Rel. Dose
6 MeV | 3 | 97.1 | 0.48917 | 3.0619 |  |  | 0.61929
9 MeV | 5 | 103 | 0.51603 | 4.9533 |  |  | 0.4669
12 MeV | 8 | 84 | 0.42105 | 8.0143 |  |  | 0.14306
16 MeV | 9.5 | 88.9 | 0.44472 | 9.5841 |  |  | 0.84092
20 MeV | 11 | 94.3 | 0.47363 | 11.0489 |  |  | 0.48861